| Scope |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Product Name: | Universal IVF Medium |  |  |  |  | Implantable: | ☐ Yes ☒ No |
| Device Description: | The product is provided in standard design (plastic bottles with a screw cap, containing a non-viscous clear (Ref# 1030) or red colored (Ref# 1031) liquid depending on content of Phenol Red), in standard packaging (cardboard boxes), with instructions for use (IFU) according to standard ART procedures provided in clear and plain language in the package insert. Further, the label on the product is marked with symbols stating to consult the IFU. The device incorporates human albumin solution (HAS), gentamicin sulphate (antibiotic), and recombinant human insulin, all considered as ancillary medicinal substances. HAS is derived from human blood constituents. Gentamicin sulphate is produced via fermentation using, among other, fish peptone. Fish peptone is the only animal-derived material used in the production of gentamicin sulphate. Intended Purpose/use: For fertilization of oocytes and culture through to cleavage stage. Can also be used for embryo transfer. |  |  |  |  |  |  |
| EU-MDR Classification: | ☐ Class IIa |  | ☐ Class IIb |  | ☒ Class III | EU Technical Documentation Number: | TD033 |
| US FDA Classification: | ☐ Class I |  | ☒ Class II |  | ☐ Class III | US Pre-Market Submission Number: | K991279 |
| Surveillance Period Date Range: |  | January 01, 2023, to December 31, 2023 |  |  |  |  |  |
| Basic UDI-DI: | 0888893TD033ZJ |  |  | Risk Management File Number: |  |  | RMF0108 |
| Associated Documents: | Document Number; Document Title. CER033R; Clinical Evaluation Report for Universal IVF Medium Version B. PMS033P; Post market surveillance plan for IVF Medium. PSUR036; Periodic Safety Update Report for IVF Medium. PMCF033; Post market surveillance plan for IVF Medium. PMCF-LRR 2023; Yearly PMCF Literature Review Report 2023. |
| Model/Catalog Numbers: | 10300060;10301010;10305060;10310060;10311010;10315060 |  |  |  |  |  |  |

Section B - Summary of Commercial Use

Universal IVF Medium has been available and used by ART clinics in Europe since 1988, when it was first introduced to the market. Apart from a long history on the market, Universal IVF Medium forms part of a group of devices characterized by common and stable designs with little evolution, and well-known safety without any safety issues in the past and with well-known clinical performance characteristics. In May 2008, a first step towards regulation of ART media as medical devices in Europe was published (Manual on Borderline and Classification in the Community Regulatory Framework for Medical Devices (version 1.1, 06-05-2008)), which was the beginning of a CE-marking process for all ART media. The same year, ORIGIO a/s (a.k.a. MediCult a/s) was one of the first companies receiving CE-mark approval of the complete ART media portfolio (March 2008). Following table summarizes the overall sales by various regions for the past 1 years starting 01 Jan 2023 and ending 31 December 2023 where 10,472 units were sold globally. Device Sales by Region and Year. Region; January 2023 - December 2023. Americas; 44. Asian Pacific; 6,018. Europe Middle East and Africa; 4,410. Total; 10,472.

Section C - Review of Customer Feedback

During the Surveillance period, data from potential sources of feedback from Users, Distributors and importers, were collected and were reviewed for feedback that would potentially affect the risk profile of Universal IVF Medium (1030 & 1031). Data was collected in accordance with BSR-MKT-005 Postmarket Surveillance and Customer Feedback, BSR-QAR-044 Complaint Handling Procedure, Clinical Evaluation Reports (Published literature, journals, and article reviews). There was no non-complaint feedback reported during the surveillance period.

Section D - Review of Product Complaint Types, Complaint Counts, and Complaint Rates

During the Surveillance period (January 01, 2023, to December 31, 2023), 10,472 units of Universal IVF Medium (1030 & 1031) were sold globally. One (1) complaint was reported during the surveillance period, therefore, the overall rate of complaints to the sales is 0.010%. Complaints vs Units Sold Year; No. of Complaints; Units Sold; Rate (%). 2023; 1; 10,472; 0.010%. Total; 1; 10,472; 0.010%. Sales and complaints are broken down regionally as shown below. Universal IVF Medium. Region; Sales; Complaints; Rate (%). Americas; 44; 0; 0.000%. Asian Pacific; 6,018; 1; 0.017%. Europe Middle East and Africa; 4,410; 0; 0.000%. Grand Total; 10,472; 1; 0.010%. The only of complaint during this surveillance period originated in the Asia Pacific Region. This complaint is associated with a missing component and no negative trends were identified.

Section E - Review of Product Complaint Rate Trending and Trend Reporting

During the surveillance period, all non-serious events that could have significant impact on the benefit-risk analysis of the subject device if their frequency or severity increase, were evaluated to determine if there had been a statistically significant increase compared to the foreseeable frequency or severity of such incidents. No Trend Reports were submitted to the competent authority for any QMS category, as there were none that indicated a statistically significant increase in severity and/or frequency.

Section F - Review of Adverse Event Reporting

Over the period between (January 01, 2023, to December 31, 2023), one (1) product complaint was collected internally, and evaluated per the Adverse Event Reporting (AER) Procedure, BSR-QAR-043. No Adverse Event Report (AER) or Medical Device Report (MDR) was identified. No new adverse event category was detected throughout the aforementioned time frame.

Section G - Review of Field Safety Corrective Actions (FSCAs)

For the reporting period of January 01, 2023 - December 31, 2023, there have been no FSCAs for Universal IVF Medium (1030 & 1031)

Section H - Review of Clinical/Technical Literature

Until date, the literature searches for Universal IVF Medium resulted in total twenty-five (25) data reports (23 publications and 2 datasets on file at CooperSurgical) on the actual device for the evaluation of clinical performance and safety after examining subject device against benchmark, or similar devices. Data extracted from the appraised literature are presented in the Clinical Evaluation Report (CER) ver. B and in the Appendices B and C of the PMCF LRR 2023. Clinical Safety There are no reports of any safety issues related to the subject device. Four (4) publications report data on safety parameters after use of the actual device, but none of them report any safety issues related to the use of Universal IVF Medium, benchmark or similar devices. Available clinical data show no trends of any negative effects of Universal IVF Medium concerning clinical safety parameters, supporting a well-recognized state of the art safety profile related to the type of device. Clinical Performance A total of twenty-three (23) data reports (21 publications and 2 datasets on file at CooperSurgical) are providing data on clinical performance related to the subject device. Based on the available clinical data, and comparing against state of the art, Universal IVF Medium is providing well-recognized state of the art performance when used as intended by professionals within ART for fertilization of oocytes and culture through to cleavage stage and for embryo transfer. In summary, there have been no changes to state-of-the-art, and the clinical safety and performance of Universal IVF Medium compliant with state-of-the-art including acknowledged KPIs, and general ART records from Europe and Australia/New Zealand, without trends of any clinical safety or performance issues. All reports of a medical device problem or harm were reported to the Product Surveillance team for assessment as potential complaints. No additional actions were determined to be needed. In summary, no new risks were identified based on the clinical evaluation, and no trends of clinical performance or safety issues were identified.

Section I - Review of External Databases and Registries

Adverse event reports involving serious incidents and recalls or relevant general notifications relating to the Universal IVF Medium products, benchmark or similar devices were identified through public databases using the following terms for further assessment (please adjust according to the device): - Product code MQL - In vitro fertilization culture medium - In vitro fertilization culture medium kit The vigilance search in MAUDE, MHRA, DAEN and SARA did not identify any reports related to the Universal IVF Medium products, benchmark or similar devices.

Section J - Review of Production Information

During the surveillance period, there were no NCMRs related to the Universal IVF Medium family of products.

Section K - Review of Corrective and Preventive Actions

During the surveillance period, there were no CAPAs related to the Universal IVF Medium family of products.

Section L - Summary of Post Market Clinical Follow-Up (PMCF)

Based on current knowledge, including clinical literature and PMS data, market history as well as the positive benefit-risk profile of the device, Universal IVF Medium is safe when used by professionals within ART and as intended for use for fertilization of oocytes and culture through to cleavage stage. It can also be used for embryo transfer. The complaint analysis in sections C and D showed no trends of performance or safety-related issues. Data supports a well-recognized safety profile related to the type of device. Additionally, long-term safety and performance have been demonstrated through clinical use for >35 years. However, after request from the British Standards Institution (BSI) to perform proactive PMCF for all Class III devices, proactive PMCF is performed on a continuous basis. As Universal IVF Medium is considered a well-established technology (WET), the time schedule for collection of data through proactive PMCF is every 4 – 5 years, to ensure new data are available and included in the CER prior to each re-certification for CE-mark. Data to be collected are clinical experience data from at least one (1) clinic where they have been using Universal IVF Medium for a minimum of 50 patients (PMCF033P, version B). Currently, we are in contact with two clinics. Data are supporting a well-recognized state of the art safety and performance profile of the device with no reports of any adverse events or negative issues with the device. As soon as all data have been collected from two clinics, a PMCF report will be finalized and referred to in the next version of the PSUR.

Section M - Summary of the Benefit-Risk Assessment

The clinical benefit of Universal IVF Medium is that it provides optimal clinical and physiological conditions to support the biological function of gametes and embryos during in vitro fertilization, in vitro culture, and embryo transfer. The product contains well-known components present in other ART media with similar intended uses and is based on known technology. Universal IVF Medium contains Ph. Eur. (European Pharmacopoeia)/USP (United States Pharmacopeia) grade HAS. Although the content of HAS is presenting a risk, the overall residual risk for Universal IVF Medium is judged to be acceptable. Pharmaceutical grade HAS is a standard component of ART media, which is also included in other CE marked media used for ART treatments. Besides HAS the medium contains other well-known components already present in other ART media with the same intended use, and which do not cause any concern regarding genotoxicity, carcinogenicity, or reproductive and developmental toxicity. The Risk Management (RM) procedure (BSR-ENG-007) has been applied to Universal IVF Medium. All RM activities have been conducted according to the procedure, and associated work instructions and forms, and the product Risk Management Report (RMF0108-RMR) has no outstanding items to be resolved. All potential risks associated with the intended purpose are minimized and acceptable as the combination of the severity and probability of the occurrence were mitigated as far as possible (AFAP). Post-production data was utilized to conduct RM for Universal IVF Medium, as described within the product PMS Plan (PMS033P) and Periodic Safety Update Report (PSUR) (PSUR033). Future PMS activities will be conducted according to BSR-QAR-042 (PMS procedure) and BSR-ENG-007 (Risk Management). This clinical evaluation did not identify any publications showing changes in state-of-the-art prohibiting use of the actual device, benchmark, or similar devices. Nor did the clinical evaluation identify any negative trends regarding performance or safety outcome after use of the actual device, benchmark, or similar devices. Finally, there were no records of Universal IVF Medium or benchmark devices in the adverse event and recall databases (USA, UK and Australia) and the product- related complaint ratio for Universal IVF Medium is 0.010%. Based on the above, Universal IVF Medium has a positive benefit-risk profile and still provides a state-of-the-art environment for fertilization and culture through to cleavage stage, as well as for embryo transfer.

Section N - Conclusion

The preceding information has been collected and reviewed as a means of assessing the overall safety and performance of the IVF Medium Product Family. The data indicates that the product has been performing acceptably and in line with expectations. The product continues to demonstrate a positive benefit-risk profile, and nothing about the information collected indicates a need for updates to the design, manufacturing process, instructions for use, labelling, risk management file, or post-market surveillance plan. When available, additional clinical data will be added to the next update of this PSUR, however, based on the Post market Data analysis, database searches and available Clinical Data, this Periodic Safety Update Report supports the overall Safety and performance of the IVF Medium products during the surveillance period