## Sheet1
Web Address | Report Number | Event Date | Event Type | Manufacturer | Date Received | Product Code | Brand Name | Device Problem | Patient Problem | PMA/PMN Number | Exemption Number | Number of Events | Event Text
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=19209743&pc=MQL | 1222779-2024-00002 | 2024-03-02 05:00 | Injury | LIFEGLOBAL GROUP, LLC | 2024-04-29 04:00 | MQL | LIFEGLOBAL MEDIA | Chemical Problem | Insufficient Information | K053552 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS IDENTIFIED THROUGH THE FERTILITY PATIENT PROGRAM (FPP) AS USING LGGG IVF PRODUCT. THE PATIENT WAS 21 WEEKS 5 DAYS PREGNANT ON (B)(6) 2024, WITH A FETUS IDENTIFIED ON ULTRASOUND AS HAVING TRANSPOSITION OF THE GREAT ARTERIES. THE COUPLE REPORTEDLY DOES NOT HAVE PERSONAL OR FAMILY HISTORY OF CARDIAC DEFECTS. NO ADDITIONAL INFORMATION IS AVAILABLE. 1222779-2024-0000002 LGGG 2024-04-0000020.Event Description: NO ADDITIONAL INFORMATION. Manufacturer Narrative: D4: MODEL #: LGGG-020 OR LGGG-050 OR LGGG-100. LOT #: 231020-0187421 OR 231020-018742 OR 231020-018743. UDI#: (B)(4). G2: FOREIGN: CANADA. H7: RECALL # 1216677-11-30-2023-001-R EVENT ID # (B)(4). DEVICE LOCATION IS UNKNOWN. ONCE THE INVESTIGATION HAS BEEN COMPLETED, A FOLLOW-UP MDR WILL BE SUBMITTED. Manufacturer Narrative: MANUFACTURING RECORD REVIEW: ALL THE DH RS RELATED TO THE LGGG RECALL THE FOLLOWING DHRS WERE REVIEWED: PRODUCT LGGG-020 LOT: 231020-018741 AND SUBASSEMBLIES (86050020 LOT: 23I020-018745 AND 8605 LOT: 231020-018744), LGGG-050 LOT 231020-0 18742 (86050050 LOT: 231020-018746 AND, 8605 LOT: 231020- 0 I 8744) AND LGGG-100 LOT 231020-018743 (86050 I 00 LOT: 231020-018747 AND 8605 LOT: 231020-0 I 8744) WERE REVIEWED AND RELATED NCMR-2023-11-00015 TO THE PROBLEM REPORTED IN THIS COMPLAINT. THE NCMR WAS RELATED TO A FAILURE OF THE MEA EMBRYONIC GROWTH AND DEVELOPMENT TEST; HOWEVER, MEA EG&D IS NOT A TEST REQUIRED AS A RELEASE REQUIREMENT. THE MEA EG&D IS A TEST THAT SOTIS THE EMB1YOS THAT DEVELOP AFTER INCUBATION IN THE TEST ARTICLE AND DOES NOT EVALUATE THE FUNCTIONALITY OF THE MEDIUM LIKE THE REGULAR MEA TEST PERFORMED ON THE TEST ARTICLE, WHICH PASSED THE TEST WITH NO PROBLEMS. FOR THAT REASON, THE PRODUCT WAS RELEASED USING THE RESULTS OF THE REGULAR MEA TEST. PRODUCT RECEIPT: THE COMPLAINT PRODUCT HAS NOT BEEN RETURNED TO COOPERSURGICAL. VISUAL EVALUATION: SAMPLES FROM THE SAME LOT NUMBER WERE INSPECTED AND AS RESULT THE COLOR APPEARANCE WERE AS EXPECTED. FUNCTIONAL EVALUATION: RETENTION SAMPLES FROM THE SAME LOT NUMBER WERE TESTED. THE MEA TEST AND PH TEST PERFORMED PASSED; HOWEVER, THE MEA EG&D TEST REQUESTED AS ADDITIONAL INFORMATION, FAILED. MOREOVER, SAMPLES RECEIVED FROM DISTRIBUTION CENTER WERE TESTED. STERILITY TEST, PH, OSMOLALITY, LAL ENDOTOXIN, GENTAMICIN AND HSA CONCENTRATION, WERE PERFORMED, AND ALL PASSED ACCORDING TO SPECIFICATIONS. PLEASE REFER TO THE QC TEST RESULTS OF THE DC SAMPLES ATTACHED TO THIS INVESTIGATION REPORT. SAMPLES RETURNED FROM CUSTOMER, RELATED TO THE SAME COMPLAINT LOT WERE TESTED AND FAILED THE MEA EG&D TEST. AS PART OF CAP A INVESTIGATION AND SINCE THE MEA EG&D FAIL , ADDITIONAL CHEMICAL TESTS WERE PERFORMED (POTASSIUM PHOSPHATE AND MAGNESIUM SULFATE) AND THE RESULTS DETERMINED AN ABSENT OF THE MAGNESIUM COMPONENT AND DOUBLE CONCENTRATION OF POTASSIUM PHOSPHATE. ROOT CAUSE ANALISYS: BASED ON THE COMPLAINT DETAILS RECORDED IN THE REPORT, IT WAS DETERMINED THROUGH OUR FPP TEAM THAT THIS PRODUCT AND LOT WAS ASSOCIATED TO THE LGGG IVF MEDIA RECALL. TO DETERMINATE THE POSSIBLE ROOT CAUSES THE A CAPA WAS OPENED. AT THE END OF THE INVESTIGATION PROCESS IT WAS DETERMINATE THAT THE POTENTIAL ROOT CAUSES WERE: 1. MATERIAL : CONTAINERS FOR STORING THE RAW MATERIAL (MAGNESIUM SULFATE AND POTASSIUM PHOSPHATE) ARE IDENTICAL IN COLOR & SHAPE. 2. METHOD: CURRENT PROCESS LACKS MEASURES TO PREVENT POTENTIAL CONFIRMATION BIAS MIX UPS WITHIN THE WEIGHING STEPS AND ADDITION OF RAW MATERIAL DURING THE FORMULATION OF THE MEDIA PRODUCT. ADDITIONALLY, AS CONTRIBUTE FACTOR IS DETERMINATE THAT THE TESTING PACKAGE WAS DESIGNED WITH THE SPECIFIC TESTING ACCORDING TO THE PRODUCT REQUIREMENTS, THAT WAS NOT CAPABLE OF DETECTING MISSING CRITICAL COMPONENTS SUCH AS MAGNESIUM SULFATE IN THE PRODUCT. THE ALLEGATIONS OF THIS COMPLAINT COULD NOT BE CONFIRMED. CORRECTIVE ACTION PLAN WILL BE HANDLED THROUGH THE CAPA. PREVENTATIVE ACTIONS BEING TAKEN: 1. IMPLEMENT A BARCODE SCANNING TECHNOLOGY TO REDUCE THE OPPORTUNITY FOR ERRORS IN THE WEIGHTING AND FORMULATION PROCESSES. CORRECTIVE ACTIONS BEING TAKEN: 1. PROCEDURE CHANGE TO ADD VISUAL VERIFICATION FROM SECOND STAFF MEMBER. VERIFICATION OF CLEAR LABELING AND SEPARATION OF INGREDIENTS DURING FORMULATION PROCESS. 2. IMPLEMENTATION OF ANALYTICAL TECHNIQUES TO IDENTIFY PRESENCE OF MAGNESIUM CONTENT WITHIN THE BULK MEDIA AFTER FORMULATION. COOPERSURGICAL WILL CONTINUE TO MONITOR THIS COMPLAINT CONDITION FOR TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18986386&pc=MQL | MW5153197 | 2024-01-27 05:00 | Injury | COOPERSURGICAL, INC. | 2024-03-26 04:00 | MQL | COOPER SURGICAL LAB MEDIA SAGE 1-STEP |  | Insufficient Information |  |  | 1 | Event Description: 21 EMBRYOS DIED DUE TO QUALITY PROBLEM IN COOPER SURGICAL LAB MEDIA USED TO DEVELOP EMBRYOS. (B)(6) 2024, COOPERSURGICAL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18935734&pc=MQL | 1216677-2024-00013 | 2024-02-26 05:00 | Injury | COOPERSURGICAL, INC. | 2024-03-19 04:00 | MQL | LIFEGLOBAL MEDIA | Chemical Problem | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT, AFTER BEING CONTACTED BY AN OPERATIONAL GENETIC COUNSELOR FROM COOPERSURGICAL, THAT SHE UNDERWENT IVF TREATMENT AND SUBSEQUENTLY BECAME PREGNANT. AT 12 WEEKS, AN ULTRASOUND ALLEGEDLY SHOWS THE FETUS TO BE MISSING A RIGHT HAND. FETUS APPEARS TO BE HEALTHY IN ALL OTHER REGARDS. ATTEMPTS HAVE BEEN MADE, BUT NO ADDITIONAL INFORMATION HAS BEEN PROVIDED. . 1216677-2024-00013 UNK LGGG IVF MEDIA (B)(4).Event Description: NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: PRODUCT LOCATION IS UNKNOWN. ONCE THE INVESTIGATION HAS BEEN COMPLETED, A FOLLOW-UP MDR WILL BE SUBMITTED. Manufacturer Narrative: D4: MODEL #: LGGG-020 OR LGGG-050 OR LGGG-100. LOT #: 231020-0187421 OR 231020-018742 OR 231020-018743. UDI#: (B)(4). H4: 20-OCT-2023 OR 16-OCT-2023 OR 31-OCT-2023. DISTRIBUTION HISTORY: THE LGGG-050 WAS MANUFACTURED AT CSI ON 20-OCT-2023 AND LGGG-020 WAS MANUFACTURED AT CSI ON 16-OCT-2023 AND LGGG-100 WAS MANUFACTURED AT CSI ON 31-OCT-2023. MANUFACTURING RECORD REVIEW: ALL THE DHRS RELATED TO THE LGGG RECALL THE FOLLOWING DHRS WERE REVIEWED: PRODUCT LGGG-020 LOT: 231020-018741 AND SUBASSEMBLIES (86050020 LOT: 231020-018745 AND 8605 LOT: 231020-018744), LGGG-050 LOT 231020-018742 (86050050 LOT: 231020-018746 AND, 8605 LOT: 231020-018744) AND LGGG-100 LOT 231020-018743 (86050100 LOT: 231020-018747 AND 8605 LOT: 231020-018744) WERE REVIEWED AND RELATED NCMR-2023-11-00015 TO THE PROBLEM REPORTED IN THIS COMPLAINT. THE NCMR WAS RELATED TO A FAILURE OF THE MEA EMBRYONIC GROWTH AND DEVELOPMENT TEST; HOWEVER, MEA EG&D IS NOT A TEST REQUIRED AS A RELEASE REQUIREMENT. THE MEA EG&D IS A TEST THAT SORTS THE EMBRYOS THAT DEVELOP AFTER INCUBATION IN THE TEST ARTICLE AND DOES NOT EVALUATE THE FUNCTIONALITY OF THE MEDIUM LIKE THE REGULAR MEA TEST PERFORMED ON THE TEST ARTICLE, WHICH PASSED THE TEST WITH NO PROBLEMS. FOR THAT REASON, THE PRODUCT WAS RELEASED USING THE RESULTS OF THE REGULAR MEA TEST. PRODUCT RECEIPT: THE COMPLAINT PRODUCT HAS NOT BEEN RETURNED TO COOPERSURGICAL. VISUAL EVALUATION: SAMPLES FROM THE SAME LOT NUMBER WERE INSPECTED AND AS RESULT THE COLOR APPEARANCE WERE AS EXPECTED. FUNCTIONAL EVALUATION: RETENTION SAMPLES FROM THE SAME LOT NUMBER WERE TESTED. THE MEA TEST AND PH TEST PERFORMED PASSED; HOWEVER, THE MEA EG&D TEST REQUESTED AS ADDITIONAL INFORMATION, FAILED. PLEASE REFER TO THE RESULTS OF THE QC TESTING OF THE RETENTION SAMPLES ATTACHED TO THIS INVESTIGATION REPORT. MOREOVER, SAMPLES RECEIVED FROM DISTRIBUTION CENTER WERE TESTED. STERILITY TEST, PH, OSMOLALITY, LAL ENDOTOXIN, GENTAMICIN AND HSA CONCENTRATION, WERE PERFORMED, AND ALL PASSED ACCORDING TO SPECIFICATIONS. PLEASE REFER TO THE QC TEST RESULTS OF THE DC SAMPLES ATTACHED TO THIS INVESTIGATION REPORT. SAMPLES RETURNED FROM CUSTOMER, RELATED TO THE SAME COMPLAINT LOT WERE TESTED AND FAILED THE MEA EG&D TEST. AS PART OF CAPA INVESTIGATION AND SINCE THE MEA EG&D FAIL, ADDITIONAL CHEMICAL TESTS WERE PERFORMED (POTASSIUM PHOSPHATE AND MAGNESIUM SULFATE) AND THE RESULTS DETERMINED AN ABSENT OF THE MAGNESIUM COMPONENT AND DOUBLE CONCENTRATION OF POTASSIUM PHOSPHATE. ROOT CAUSE ANALISYS: BASED ON THE COMPLAINT DETAILS RECORDED IN THE REPORT, IT WAS DETERMINED THROUGH OUR FPP TEAM THAT THIS PRODUCT AND LOT WAS ASSOCIATED TO THE LGGG IVF MEDIA RECALL. TO DETERMINATE THE POSSIBLE ROOT CAUSES A CAPA WAS OPENED. AT THE END OF THE INVESTIGATION PROCESS IT WAS DETERMINATE THAT THE POTENTIAL ROOT CAUSES WERE: I. MATERIAL: CONTAINERS FOR STORING THE RAW MATERIAL (MAGNESIUM SULFATE AND POTASSIUM PHOSPHATE) ARE IDENTICAL IN COLOR & SHAPE. 2. METHOD: CURRENT PROCESS LACKS MEASURES TO PREVENT POTENTIAL CONFIRMATION BIAS MIX UPS WITHIN THE WEIGHING STEPS AND ADDITION OF RAW MATERIAL DURING THE FORMULATION OF THE MEDIA PRODUCT. ADDITIONALLY, AS CONTRIBUTE FACTOR IS DETERMINATE THAT THE TESTING PACKAGE WAS DESIGNED WITH THE SPECIFIC TESTING ACCORDING TO THE PRODUCT REQUIREMENTS, THAT WAS NOT CAPABLE OF DETECTING MISSING CRITICAL COMPONENTS SUCH AS MAGNESIUM SULFATE IN THE PRODUCT. THE ALLEGATIONS OF THIS COMPLAINT COULD NOT BE CONFIRMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18815759&pc=MQL | 18815759 | 2023-11-27 05:00 | Malfunction | LIFEGLOBAL GROUP, LLC | 2024-03-01 05:00 | MQL | LIFEGLOBAL® MEDIA |  | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: CULTURE MEDIA DEFINITELY DID NOT SUPPORT EMBRYO DEVELOPMENT TO THE BLASTOCYST STAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18704518&pc=MQL | MW5151585 | 2023-11-20 05:00 | Injury | COOPERSURGICAL INC. | 2024-02-13 05:00 | MQL | COOPER SURGICAL GLOBAL MEDIA |  |  |  |  | 1 | Event Description: ON (B)(6) 2023 I HAD AN IVF(IN VITRO FERTILISATION) EGG RETRIEVAL DONE WHICH YIELDED 12 EGGS, 8 OF WHICH FERTILIZED. ALL EMBRYOS ARRESTED DEVELOPMENT ON DAY 3 AND FAILED TO MAKE IT TO BLASTOCYST STAGE. COOPER SURGICAL GLOBAL MEDIA USED WITH LOT #230922-017934.